## Slide 1: ICG Sensor
Group 17
Daniel Arfstrom, Tuyen Do, Chris McCord, Stephen Wilkes
Sponsor: Dr. Thomas Looke (Anesthesiologist/EE)

## Slide 2: Indocyanine Green
Medical Dye used for its fluorescent properties.
Comes in powdered form, mixed with water to create an injectable dye.

## Slide 3: SPY Elite System ( SPYE)
Used during surgeries, typically to pinpoint areas with little to no blood flow
Room must be darkened to filter out external light
Extremely costly
Specific to only certain surgeries
Only usable in operating room

## Slide 4: Goals and Objectives
Safe
Compact
Easy to use
Easily read output

## Slide 5: What is the Purpose of the Final Product?
“There are two uses that I have in mind for this device. 1. The steady state value of fluorescence (shortly after injection prior to the dye being metabolized) should give an indicator of dye concentration which can be used to calculate blood volume. 2. The dynamic response, plot of fluorescence vs. time, is related to cardiac output. So we should have an indicator of both blood volume and cardiac output with this dye fluorescence time plot.” -Dr. Looke

## Slide 6: System I/O
Emitter
Medium w/ ICG
Photodiode
Micro-Controller
Display
Data Processing

## Slide 7
ICG Sensor Block Diagram
Indocyanine Green
Collector
Emitter
MCU
Touchscreen LCD
Battery
Fluorescence
Near Infrared Light
Analog signal
5V
Data points
Button Presses
PSU

## Slide 8: Properties of ICG
Indocyanine Green possesses useful optical properties.
Absorbs light of a specific range of frequencies, steps it down in energy, and re-releases it (as fluorescence).
ICG's emission frequency is altered slightly in blood plasma, as it binds to proteins.
Ideal measure wavelength ~830 nm.

## Slide 9: Properties of ICG
Process called "Quenching" makes fluorescence decrease with concentration after a certain concentration is reached.
Our measurable range then is from 8 to 100 micrograms per milliliter (before quenching takes effect).

## Slide 10: Photodiode Substrate Choices
Silicon photodiodes have an optimal wavelength range around 830 nm and are more cost effective.
Substrate | Wavelength Range | Typical Pricing
Si | 350-1100 nm | $13-$100
Ge | 800-1800 nm | $130-$430
InGaAs | 800-1800 nm | $130-$260

## Slide 11: Si Photodiode Choices
Item # | Range (nm) | Active Area | Dark Current | Price
FDS010 | 200-1100 | .82 mm2 | .3 nA@10V | $42.10
FDS10X10 | 340-1100 | 100 mm2 | 200 pA@5V | $100.00
FDS100 | 350-1100 | 13 mm2 | 35 pA@5V | $73.50
FDs02 | 400-1100 | .049 mm2 | 35 pA@5V | $73.50
FDS1010 | 400-1100 | 100 mm2 | 1.05 nA@5V | $48.80

## Slide 12: Si Photodiode Selection

## Slide 13: Filter Choices
The same manufacturer had a line of 10 nm FWHM bandpass filters with 1 inch diameter
Shown left is the transmission graph of the selected 830nm-centered bandpass filter.

## Slide 14: Sensing Range

## Slide 15: Emitter Selection
Simply needed excitation source within absorption range.
Excitation light should be a broader range than collection, so that it is made sure that ICG fluoresces within the pass band.

## Slide 16: Emitter Selection
Thorlabs' LED780E
780 nm centered with FWHM 30 nm
190 mW max power dissipation
100 mA max DC forward current
1.75 V typical forward voltage

## Slide 17: Emission/Collection Ranges

## Slide 18: Alternate Filter for Testing
Since real blood is not a testing option, the focus must shift to the properties of ICG for an aqueous solution.
The peak fluorescence for such a situation is 800 to 810 nm, instead of 830 for blood plasma.
This project implements a Thorlabs 810 nm version of the previously mentioned filter.

## Slide 19: Complete Sensor Circuitry

## Slide 20: PCB Schematic

## Slide 21: PCB Layout

## Slide 22: Output
Primary display:
Main body: Graphical display of collected data.
ICG Level indicator:
Indicator for user to note critically low level of ICG
.
Snapshot display:
Main body: Various Statistics and most recent 12 data points.

## Slide 24: Comparison of Displays
Display | Vin | Size | Resolution | PPI | Color | Backlight | I/O | Touchscreen | Price
ILI9325 | 3-5V | 2.8” | 320x240 | 142 | 18 bit | Yes | 12 lines | Resistive | $40.00
GDM12864HLCM | 4.5-5.5V | 2.4” | 128x64 | 59 | 1 bit | Yes | 10 lines | None | $19.95
μLCD-32PTU | 4-5.5V | 3.2” | 320x240 | 125 | 16 bit | Yes | 13 lines | Resistive | $84.95
Preferably be powered by 5V
~3” (diagonal) in size
High PPI for time-magnitude graph
Color
Backlight for low light environments

## Slide 25: Display Specifications
Adafruit 2.8” TFT LCD (ILI9325)
  3-5V
  8 Digital, 4 Analog Control Lines
  2.8" TFT LCD
  320x240 Resolution
  18 Bit Color (262,000 Colors)
  LED Backlight
  Resistive Touchscreen
  RGB, SPI Interfaces

## Slide 26: Comparison of Microcontrollers
ATmega328P-PU
1.8-5.5V
20 MHz
32kB Flash
2kB RAM
23 GPIO pins
8ch 10-bit ADC
I2C, SPI, UART Interfaces
MSP430
1.8-3.6V
16 MHz
16kB Flash
512B RAM
16 GPIO pins
8ch 10-bit ADC
I2C, SPI, UART Interfaces

## Slide 27: Originally…
We chose the MSP430G2553 due to
  Our familiarity with the microcontroller
  Sufficient GPIO pins
  Launchpads readily available

## Slide 28: ATmega328P
Advantages over the MSP430
  Simplified power requirements; no longer needed a 3V regulator.
  2x the flash memory; software currently exceeds 20 kB.
  4x the RAM will allow the software to do more things at once.
  Extra pins for expandability.
  More abundant software libraries.

## Slide 29: Software States
Collection
Storing
Drawing
Touch Listener

## Slide 30: Data Storage Implementation
Circular Linked List
  Limited space means old data will be purged for new incoming data (FIFO).
  Two pointers will be needed: head and tail.
Running Times
  Insertion O(1) – we will only be inserting at the end of the list.
  Search O(n)
  Deletion O(n)

## Slide 31: Power System
Rechargeable Battery Pack | 
Type | LIPO
Cells | 2
Voltage | 7.4 V
Capacity | 2200mAh
Max 846A Linear IC

## Slide 32: Why LIPO?
1.Li-ion has slower discharge but it is not a problem due to consistent use.
2.High per cell voltage.
3.High capacity.
4.Switching is easy!
(Picture courtesy of Battery University's Isidor Buchmann)

## Slide 33: Voltage Regulation
Voltage regulator: LM7805
All components run off 5V except for small red 3V LED; LED 780E, photodiode, LCD, microcontroller, and reference voltage. 100 ohm resistor in series with the LED will take care of voltage difference.
Input (V) | Output (V)
11 | 5.01
10 | 5.01
9 | 5.01
8 | 5.01
7 | 5.01
6 | 4.99
5 | 4.06

## Slide 34: Charging Circuit Block Diagram
Transformer
Resettable Fuse
Rectifier
LIPO IC
Resettable Fuse
Wall Socket
Battery

## Slide 35: Charging IC
MAX846EVKIT+

## Slide 36: Charger Casing
Connection from the wall socket, to charger circuit, to the main housing unit.
Made of Aluminum.
Simple rectangular/box casing.
7 in.
5.5in.

## Slide 37: Main Housing Unit
Contains the rechargeable battery, MCU, LCD, and PCB
Flips open where one side houses the battery and the other houses the PCB, MCU, and LCD
Easy access to battery makes replacement simple
3mm hole for red LED "on light"
Mechanical on/off switch
Reset switch
Wires running from the main housing to the sensor are twisted in a helical formation to limit noise and are kept together using heat shrink tubing
3.5 in.
8 in.

## Slide 38: Sensor Casing
Contains the collector, emitter, and filter.
Oval in shape
Easy to hold with flat side corners
Must be constantly held over test area by user.
4 in.
3 in.

## Slide 39: Bottom View of Sensor
Collector has a < 1 in. in diameter hole for filter to reside
Physical separation between collector and emitter by bent aluminum plate
Wire runs through the hole in separation to collector
Metalized Mylar in the walls of the entire closure.
Terminal block for convenience and aesthetics to hold wires

## Slide 40: Aluminum Foil Reflectivity
Slight drop at 800 nm

## Slide 41: Testing Medium, Synthetic Blood
The ideal testing environment would be using a human patient.
Problems with legality.
Next option, Synthetic blood.
Needed enough for thorough testing.
Issue with customs; importing 5 liters of "blood".

## Slide 42: Testing methods
Testing materials will be covered from external light to avoid false ICG fluorescence and false sensor readings.
Clear plastic containers and bags will be used to contain the synthetic blood and ICG.
Were going to use Synthetic skin and muscle, but determined it would not be necessary for a proof of concept.
Magnitude of fluorescence of ICG is understood to be directly related to ICG concentration in the testing medium (Synthetic blood).

## Slide 43: Primary Tests
Three Tests:
Responsiveness Test
Infusion Test
Decay Test
Responsiveness Test:
Testing to ensure collector responds to an LED that emits a readable signal.
Turn on sensor and start recording.
Introduce the powered LED to the collector.
Verify the device indicates an increased signal magnitude.

## Slide 44: Primary Tests (Cont’d)
Infusion Test:
Start with a clean volume of testing medium.
Turn on sensor and start recording.
Slowly introduce a sample of ICG and verify that the device indicates a trend of increasing magnitude of fluorescence.
Decay Test:
Start with a volume of testing medium with a predetermined quantity of ICG .
Turn on sensor and start recording.
Slowly introduce a quantity of clean testing medium. Verify the device indicates a decreasing magnitude of fluorescence.

## Slide 45: Budget
Part | Cost | Quantity | Shipping | Notes | Total
Atmel ATMEGA328P-PU | $0.00 | 3 | $0.00 | Samples | $0.00
Adafruit 2.8” TFT LCD Display | $40.00 | 1 | $3.99 | N/A | $43.99
Synthetic Blood (5 liters) | $287.67 | 1 | $57.53 | N/A | $345.20
LIPO Rechargeable Battery | $33.34 | 1 | $0.00 | N/A | $33.34
PCB | $23.75 | 8 | $0.00 | N/A | $190
Charging Unit | $255.15 | 1 | $0.00 | N/A | $255.15
Wire/Cable | $3.94 | 1 | $0.00 | N/A | $3.94
FDS1010 Photodiode | $48.80 | 1 | $8.58 | N/A | $57.38
830nm Bandpass Filter | $84.67 | 1 | $0.00 | N/A | $84.67
LED780E x5 | $26.00 | 1 | $0.00 | N/A | $26.00
Misc. | $19.50 | 1 | $0 | N/A | $19.50
Total |  |  |  |  | $1059.17

## Slide 46: Distribution of Work
 | ICG | Sensor | MCU | LCD | PSU | Coding | Casing | Testing
Daniel |  |  | X | X |  | X |  | X
Tuyen |  |  | X | X |  | X |  | X
Chris | X |  |  |  | X |  | X | X
Stephen | X | X |  |  |  |  | X | X

## Slide 47: Issues
Touchscreen not functioning properly.
ICG:
  Provided in limited quantities.
  Scale is not very precise.
  Not performing as expected.
  The possibility exists that emitter light is not sufficient.
Testing
  Noise
  Precision
  Compatibility between synthetic blood and real blood with ICG

## Slide 48: Questions
Thank you for your attention, feel free to ask any questions.